## Slide 1
1-D DIC IMPLEMENTATION ON UNIVERSAL TESTING MACHINE
A camera-instrumented desktop tensile tester — built, validated against a commercial lab, and documented.
Adithya Sivakumar
adithya.sivakumar@ju.se
Supervisors: Johan Jansson · Mirza Cenanovic
SEPTEMBER 2026
01

## Slide 2
AGENDA
Today, in six parts
PPD-UTM · FINAL PRESENTATION
02
01
DIC, in one picture
what the rig measures, and why no sensor touches the part
02
What was built
hardware, software — and 19 things it does on its own
03
Live demo
a specimen pulled to fracture, live on the rig
04
Validation
the same material through a commercial lab — and how close we land
05
AI in the loop
how this was built by one engineer and a coding agent
06
What comes next
future work, and one number for fun
Appendix afterwards: Otsu thresholding · PLA vs PETG vs TPU · rig-to-rig noise.

## Slide 3
01 · BACKGROUND
DIC — the camera is the strain gauge
BACKGROUND
03
DIC — Digital Image Correlation — measures how a part deforms by comparing photographs of its surface while it is loaded; here in its simplest two-dot form.
No contact
nothing touches the specimen — no clip-on gauge to knock off at fracture
No calibration
strain is a ratio of two pixel distances — mm per pixel cancels out
Every frame
35 measurements per second, live on screen while the test runs

## Slide 4
02 · THE BUILD
One rig, one application — built together
THE BUILD
04
the PPD-UTM
HARDWARE
▪ twin ball-screw frame, 4.5 kN load cell
▪ industrial camera + LED panel — the DIC eye
▪ printed grips, emergency stop, enclosure
SOFTWARE
▪ one PyQt6 application runs the whole test
▪ live DIC strain at 35 fps, six closed-loop protocols
▪ every run leaves a CSV + video + PDF report
▪ post-processor re-measures any recorded video
28
specimens pulled
3
materials
19
smart features
~44 k
lines of Python
204
evidence pages
numbers as of Sep 2026

## Slide 5
02 · THE BUILD
Nineteen things the rig does on its own
THE BUILD
05
SF1
DIC health HUD
SF2
Prepare specimen
SF3
Settings / recipes
SF4
1-click PDF report
SF5
Auto-stop at fracture
SF6
Strain-rate control
SF7
Stall guard
SF8
Release load
SF9
6 test protocols
SF10
Auto-preload
SF11
Auto-metadata link
SF12
DIC auto-calibrate
SF13
Guided wizard
SF14
Test registry
SF15
Dead-DIC guard
SF16
Live Px₀ overlay
SF17
Video + stills capture
SF18
DIC post-processing
SF19
Noise capture
each of these has its own evidence page in the project deck

## Slide 6
02 · THE BUILD
…and the three worth two minutes each
THE BUILD
06
SF1
DIC health HUD
SF2
Prepare specimen
SF3
Settings / recipes
SF4
1-click PDF report
SF5
Auto-stop at fracture
sees the load collapse and halts the motor itself — no hand on the stop button
SF6
Strain-rate control
closes the loop on measured strain rate, not motor speed — constant dε/dt
SF7
Stall guard
SF8
Release load
SF9
6 test protocols
SF10
Auto-preload
SF11
Auto-metadata link
SF12
DIC auto-calibrate
SF13
Guided wizard
SF14
Test registry
SF15
Dead-DIC guard
SF16
Live Px₀ overlay
SF17
Video + stills capture
SF18
DIC post-processing
re-measures ANY recorded video — it is how the commercial lab's test got audited
SF19
Noise capture
SF5 protects the specimen · SF6 makes tests comparable · SF18 made the validation possible

## Slide 7
03 · LIVE DEMO
Three materials, three behaviours — then the real thing
LIVE DEMO
07
PLA — S13
PLA
black PLA · fractures at ~4 % strain
PETG — S30
PETG
glossier, more ductile · ~8 % at fracture
TPU — S35
TPU
rubber-like · leaves the frame before it fractures
15-second timelapses of real pulls on this rig (click to play) — then we switch to the live software and break one for real.

## Slide 8
04 · VALIDATION
The same material, through a commercial lab
VALIDATION
08
MOT LAB PHOTO
drops in here — machine + XT-205 in view
THE REFERENCE
▪ MOT materials lab — commercial tensile frame
▪ XT-205 video extensometer: a camera that tracks two painted marks — the same principle as our DIC, in certified form
▪ two specimens from our own print batch: XT205-S1 and XT205-S2
TEST CONDITIONS · XT205-S2
45 mm gauge
2 mm/min
300 N preload
▪ same stress-strain quantities, computed two independent ways
▪ and their raw video, re-measured by OUR post-processor (SF18)

## Slide 9
04 · VALIDATION
Two rigs, one material — the curves land together
VALIDATION
09
S34 (PPD-UTM) / XT205-S2 (MOT)
UTS
48.34 / 44.62 MPa
−7.7 %
Yield σy
39.23 / 39.61 MPa
+1.0 %
Modulus E
3.45 / 3.23 GPa
−6.4 %
Fracture εf
5.25 / 5.21 %
−0.8 %
▪ strain agrees to 1 % — the DIC is doing its job
▪ stress gap is real but small — different specimens, different day, 3× pull speed
▪ shape of the whole curve matches to fracture

## Slide 10
04 · VALIDATION
Same video, two calculations — one line
VALIDATION
10
0.99928
scale — ours vs theirs
0.072 % from perfect
−55 µε
constant offset
= 0.07 px on a 1255 px gauge
0.9999964
R², 3 153 matched instants
one video, two independent codes
The MOT lab's own fracture video, fed through our post-processor: the two strain calculations are the same measurement.

## Slide 11
04 · VALIDATION
Where the remaining gap comes from
VALIDATION
11
Pull speed
measured
▪ their pull: 2 mm/min · ours: ≈3× faster
▪ but at the GAUGE the strain-rate ratio is only 1.19×
▪ → predicts under 1 % of the stress gap — a real effect, not the story
Grips & seating
measured
▪ only 21–45 % of crosshead travel reaches the gauge
▪ our four runs cluster at 21 % and 33 % — a joint that either seats or doesn't
▪ → affects travel-based numbers, never the DIC strain
Specimen & moisture
study pending
▪ different print batches, weeks apart, unconditioned storage
▪ PLA is hygroscopic — literature review queued
▪ → explanation lands here once the study is done
The honest ranking: seating and specimen dominate; speed is measured and small.

## Slide 12
05 · PROCESS
Built by one engineer and a coding agent
PROCESS
12
intent
“the rig must stop itself at fracture”
↓
prompt
precise ask + constraints + how to verify
↓
Claude Code
writes app code, tests, analysis
↓
verify on rig
real pulls; failures go straight back in
↓
evidence
every claim becomes a measured deck page
348
commits · 94 % co-authored with the agent
9 weeks
from first DIC frame to a validated rig — 314 commits
~44 k
lines of Python: app, analysis, this deck's generator
204
evidence pages, regenerated from raw data on every build
▪ the prompt is the engineering: what to build, what would falsify it, how to prove it ran
▪ the agent measures instead of assuming — it renders every page and looks at it

## Slide 13
06 · OUTLOOK
Future work — named, sized, and how
OUTLOOK
13
2D-DIC strain field
full-field strain over the speckled gauge, not just two dots
HOW
offline first, in the existing post-processor; needs lower exposure — the white body currently saturates the sensor
Correlation, not binary
sub-pixel tracking that keeps the grey levels
HOW
measured gain: 1.1–2.1× less jitter; add as a post-processing option before touching the live loop
Feed simulation models
measured strain fields calibrating FE material cards
HOW
export the field per frame; compare predicted vs measured maps on the same geometry
Poisson & true stress (FW1)
four markers → lateral strain, ν, Cauchy stress
HOW
maths written and self-tested; blocked on optics — needs a gauge-zoomed second camera or a wider specimen
Grip-seating experiment (FW2)
why gauge share clusters at 21 % vs 33 %
HOW
one specimen, re-mounted several times, elastic pulls only — spread survives → seating
Rate-matched MOT rerun
close the last confound in the lab comparison
HOW
pull at 5.01 mm/min (rate-matched) and 2.00 mm/min (speed-matched) and re-compare

## Slide 14
06 · OUTLOOK
One number for fun — the carbon footprint
OUTLOOK
14
CARBON-FOOTPRINT NUMBERS LAND HERE
the calculation is queued: filament + rig power + compute for the AI collaboration
spoiler: the printed specimens are not the biggest line item.

## Slide 15
Thank you.
Questions — or come break a specimen with me.
Adithya Sivakumar
adithya.sivakumar@ju.se
15

## Slide 16
APPENDIX
For the curious
APPENDIX
16
A1
Otsu thresholding
how the camera decides what is a marker
A2
PLA vs PETG vs TPU
our curves against what literature expects
A3
Noise, rig against rig
how quiet each instrument is, measured at rest

## Slide 17
APPENDIX · A1
Otsu — how the camera picks its threshold
APPENDIX
17
▪ WHY it exists
▪ every frame must be split into “marker” and “not marker” — that takes one brightness number
▪ WHAT it does
▪ tries all 254 cuts, keeps the one giving the two tidiest groups — lands in the histogram's empty valley
▪ THE CATCH
▪ it follows the SCENE, not the markers — so the rig scores it against fixed values (auto-calibrate) instead of trusting it blindly
Otsu (1979), IEEE Trans. SMC 9(1):62–66 · measured here on 120 frames of specimen S13.

## Slide 18
APPENDIX · A2
Three materials — measured, against expectations
APPENDIX
18
▪ ordering exactly as published
▪ TPU never fractured — it left the frame at 19 % strain
published ranges: filament TDS + printed-part literature, as compiled in the project's material notes.

## Slide 19
APPENDIX · A3
How quiet is each instrument?
APPENDIX
19
▪ both rigs sit in the tens of microstrain
▪ our floor is set by pixel quantisation: noise ∝ 1/Px₀
▪ predicted 1.78× between our two gauge lengths — measured 1.78×
▪ so the noise is understood, not just small