## Slide 1
1-D DIC IMPLEMENTATION ON UNIVERSAL TESTING MACHINE
A camera-instrumented desktop tensile tester — built, tested, validated, and documented.
Adithya Sivakumar
adithya.sivakumar@ju.se
Supervisors: Johan Jansson · Mirza Cenanovic
SEPTEMBER 2026
01

## Slide 2
AGENDA
Content, in six parts
02
01
DIC, in one picture
what the rig measures, and why no sensor touches the part
02
What was built
hardware, software — and 20 things it does on its own
03
Live demo
a specimen pulled to fracture, live on the rig
04
Validation
the same material through a MOT lab test — and how close we land
05
AI in the loop
how this was built by one engineer and a coding agent
06
Future work
What comes next and interesting numbers
Appendix: Thresholding · PLA vs PETG vs TPU · rig-to-rig noise.

## Slide 3
01 · BACKGROUND
Where DIC is used — and where this project sits
BACKGROUND
03
DIC — Digital Image Correlation — measures how a surface deforms by comparing photographs of it under load: no contact, no wires, just pixels tracked frame to frame.
Material testing & characterization
stress–strain curves, E, σy, UTS, εf — strain measured without touching the specimen.
THIS PROJECT
Material testing & characterization
stress–strain curves, E, σy, UTS, εf — strain measured without touching the specimen. THIS PROJECT: a desktop tensile rig whose only extensometer is a camera.
Manufacturing quality control
warpage, springback and assembly distortion checked on the line — parts against CAD, contactless and in minutes
Aerospace structural monitoring
wing bend and fuselage-skin strain in full-scale tests — thousands of virtual gauges from a pair of cameras
Fracture mechanics & crack propagation
strain concentration at the crack tip, crack speed and direction — our future-work item 4 builds exactly this on the same rig
Biomechanics & medical devices
tissue, implants and prostheses under load — contactless is the only option when the specimen must stay untouched
Thermal stress analysis
expansion and warpage under heating — where a glued strain gauge would cook, the camera just watches
…and more: FEA model validation, civil-structure monitoring, electronics reliability, high-speed impact. One method, many arenas — the next slides are our corner: the tensile test.

## Slide 4
01 · BACKGROUND
DIC — the camera is the strain gauge
BACKGROUND
04
How WE use DIC — the simplest form there is: two painted dots on the gauge; the camera tracks their separation L every frame, and strain is just (L − L₀) / L₀.
Contactless
No strain gauge glued on, no clip-on extensometer — nothing touches the specimen
Pixels become strain
Each frame, the marker spacing L is measured in pixels and compared with the frozen L₀ — strain = (L − L₀) / L₀, no millimetres needed
Live stress–strain
Load cell gives force, the camera gives strain — the stress–strain curve draws itself live during the pull
Notes: glued = strain gauge, clipped = extensometer, camera = video extensometer (XT-205, and our rig)

## Slide 5
01 · BACKGROUND
DIC — the camera is the strain gauge
BACKGROUND
05
DIC — Digital Image Correlation — measures how a part deforms by comparing photographs of its surface while it is loaded; here in its simplest two-dot form.
the clip-on extensometer — the principle, and a commercial one (Zwick/Roell): knife edges TOUCH the specimen
our rig — the camera, touching nothing
Contactless
No strain gauge glued on, no clip-on extensometer — nothing touches the specimen
Pixels become strain
Each frame, the marker spacing L is measured in pixels and compared with the frozen L₀ — strain = (L − L₀) / L₀, no millimetres needed
Live stress–strain
Load cell gives force, the camera gives strain — the stress–strain curve draws itself live during the pull

## Slide 6
01 · BACKGROUND
DIC — the camera is the strain gauge
BACKGROUND
06
DIC — Digital Image Correlation — measures how a part deforms by comparing photographs of its surface while it is loaded; here in its simplest two-dot form.
Contactless
No strain gauge glued on, no clip-on extensometer — nothing touches the specimen
Pixels become strain
Each frame, the marker spacing L is measured in pixels and compared with the frozen L₀ — strain = (L − L₀) / L₀, no millimetres needed
Live stress–strain
Load cell gives force, the camera gives strain — the stress–strain curve draws itself live during the pull

## Slide 7
01 · BACKGROUND
DIC — the camera is the strain gauge
BACKGROUND
07
DIC — Digital Image Correlation — measures how a part deforms by comparing photographs of its surface while it is loaded; here in its simplest two-dot form.
S26 · PLA — the curve exactly as the rig draws it live, one point per frame
Contactless
No strain gauge glued on, no clip-on extensometer — nothing touches the specimen
Pixels become strain
Each frame, the marker spacing L is measured in pixels and compared with the frozen L₀ — strain = (L − L₀) / L₀, no millimetres needed
Live stress–strain
Load cell gives force, the camera gives strain — the stress–strain curve draws itself live during the pull

## Slide 8
02 · THE BUILD
DIC rig and Application built together
THE BUILD
08
the PPD-UTM DIC
HARDWARE
▪ Aluminium-profile frame — twin ball screws, brackets, load holders
▪ 4.5 kN load cell, 3D-printed grips
▪ Drive motors + their control electronics, emergency stops
SOFTWARE — BUILT IN-HOUSE
▪ One PyQt6 application runs the whole test — written line by line in this project
▪ OpenCV does the seeing: thresholding, blob detection, sub-pixel centroids
▪ Live DIC strain at ≈20 fps (measured — the exposure sets it), six closed-loop protocols
▪ 20 smart features, each automated behind a single click
▪ Basler camera · LED panels · light enclosure — the DIC eye, and the hardware I worked on most

## Slide 9
02 · THE BUILD
20 Smart Features - rig does on its own
THE BUILD
09
SF1
DIC health HUD
SF2
Prepare specimen
SF3
Settings / recipes
SF4
1-click PDF report
SF5
Auto-stop at fracture
SF6
Strain-rate control
SF7
Stall guard
SF8
Release load
SF9
Advanced test modes
SF10
Auto-preload
SF11
Auto-metadata link
SF12
DIC auto-calibrate
SF13
Guided wizard
SF14
Test registry
SF15
Dead-DIC guard
SF16
Live Px₀ overlay
SF17
Video + stills capture
SF18
DIC post-processing
SF19
Noise capture
SF20
Manual blob selection
each of these has its own evidence page in the final project deck

## Slide 10
02 · THE BUILD
3 of my best picks for the smart features
THE BUILD
10
SF1
DIC health HUD
SF2
Prepare specimen
SF3
Settings / recipes
SF4
1-click PDF report
SF6
Strain-rate control
SF7
Stall guard
SF8
Release load
SF10
Auto-preload
SF11
Auto-metadata link
SF12
DIC auto-calibrate
SF13
Guided wizard
SF14
Test registry
SF15
Dead-DIC guard
SF16
Live Px₀ overlay
SF17
Video + stills capture
SF19
Noise capture
SF20
Manual blob selection
SF5
Auto-stop at fracture
sees the load collapse and halts the motor itself — no hand on the stop button
SF9
Advanced test modes
6 closed-loop protocols in one tab — the tests beyond a plain pull to fracture
SF18
DIC post-processing
re-measures ANY recorded video — it is how the commercial lab's test got audited
SF5 protects the specimen · SF9 runs the six closed-loop protocols · SF18 made the validation possible

## Slide 11
03 · LIVE DEMO
PLA snaps, PETG draws, TPU stretches — then live demo
LIVE DEMO
11
PLA — S26
PLA
quasi-brittle — snaps in one clean break at ~5 % strain
PETG — S30
PETG
glossier, more ductile — it draws, then tears
TPU — S35
TPU
rubber-like · leaves the frame before it fractures
15-second timelapses of real pulls on this rig — then we switch to the live software and break one for real.

## Slide 12
04 · VALIDATION
The same specimen, tested at MOT lab
VALIDATION
12
the MOT lab: Zwick/Roell Z100, XT-205 video extensometer on our specimen
THE REFERENCE
▪ MOT lab — commercial tensile frame
▪ XT-205 video extensometer: a camera that tracks two painted marks — the same principle as our DIC, in certified form
▪ Two specimens from our own print batch: XT205-S1 and XT205-S2
TEST CONDITIONS · XT205-S2
45 mm gauge length
2 mm/min
300 N preload
TWO ANALYSES – XT205 at MOT against PPD-UTM DIC rig
▪ Analysis 1 — instrument vs instrument: the XT-205's certified strain readout against our rig's DIC on matched specimens; E, σy, UTS and ε_f each computed twice, independently
▪ Analysis 2 — our calculation method on XT205 at MOT test data and video: the session's raw XT-205 video re-measured frame-by-frame by our DIC post-processor (SF18) on the lab’s machine’s footage

## Slide 13
04 · VALIDATION
Analysis 1 : Two rigs, two specimens, one material
VALIDATION
13
S34 (PPD-UTM DIC) / XT205-S2 (MOT)
UTS
48.34 / 44.62 MPa
−7.7 %
Yield σy
39.23 / 39.61 MPa
+1.0 %
Modulus E
3.45 / 3.23 GPa
−6.4 %
Fracture εf
5.25 / 5.21 %
−0.8 %
▪ UTS and E gaps (6–8 %) carry a 3× speed difference plus normal specimen scatter (CV ≈ 2.5 %) — within family, and the whole curve shape matches
▪ Yield within 1 % — stress scale agrees where the curve is stable
▪ Fracture strain 5.25 vs 5.21 % — two different optical instruments, 0.8 % apart
▪ Analysis 2 closes the loop: their own video through our post-processor lands on the same line (two slides on)

## Slide 14
04 · VALIDATION
Analysis 1 : why the offset?
VALIDATION
14
Grips & seating
measured
▪ 50 % of the EARLY crosshead travel (under 1% strain) is eaten by slack in our printed grips before the gauge feels anything
▪ our four runs cluster at 21 % of travel reaching the gauge — a joint that either seats or doesn't
▪ → costs only EARLY crosshead travel and bends the curve's toe — the DIC strain never sees it: the camera watches the gauge, not the crosshead
Pull speed
measured
▪ Travel speed we INPUT is the CROSSHEAD's, in mm/sec (ours 0.1 mm/sec) — an open loop: no feedback tracks whether the specimen is stretched at that speed
▪ → Strain-rate moves the SLOPE (E) more than the UTS — and predicts under 1 % of the stress gap
Specimen & moisture
studied
▪ Different print batches, weeks apart, unconditioned storage
▪ even within ONE batch, UTS scatters print-to-print (CV ≈ 2.5 %) — the batch effects sit on top of that
▪ literature: 24 h of humid filament storage already costs PLA ≈8.5 % UTS; 150 h ≈11 % (Hamid 2022)
Hamid et al. 2022 · doi 10.1007/978-981-16-8954-3_47

## Slide 15
04 · VALIDATION
Analysis 2 : Same video, Same specimen, two calculation methods
VALIDATION
15
+0.08 %
fracture strain apart
5.21 (XT-205) vs 5.22 (PPD-UTM DIC), on the same footage
−55 µε
constant difference, flat to fracture
= 0.07 px on a 1255 px gauge
3.21–3.24 vs 3.23
Young's modulus E (GPa), elastic region — PPD-UTM DIC fits on the XT-205 footage vs the XT-205 machine
fixed-window fit 3.21 · steepest-window 3.24 · machine 3.23 — the SLOPE agrees, not just the endpoints

## Slide 16
05 · PROCESS
AI in the loop - Built by one engineer and a coding agent
PROCESS
16
intent
“the rig must stop itself at fracture”
prompt
precise ask + constraints + how to verify
Claude Code
writes app code, tests, analysis
verify on rig
real pulls; failures go straight back in
evidence
every claim becomes a measured deck page
370
commits · 94 % co-authored with the agent
~26 k
lines of Python in the rig software itself — app, DIC engine, analysis, tests (deck-builder scripts not counted)
853
prompts, measured from the session log (28 May → 6 Sep) — about two per shipped commit
≈7×
the leverage: 195 h hands-on with the agent over 6 months vs ≈1 300 h estimated solo (26 k lines of hardware-coupled GUI + vision code at ~20 finished lines/hour)
▪ the prompt is the engineering: what to build, what would falsify it, how to prove it ran
▪ the agent measures instead of assuming — it renders every page and looks at it

## Slide 17
06 · OUTLOOK
Future work
OUTLOOK
17
Poisson & true stress
four markers → lateral strain, ν, Cauchy stress
HOW
maths written and self-tested; blocked on optics — needs a gauge-zoomed second camera or a wider specimen
2D-DIC strain field
full-field strain over the speckled gauge, not just two dots
HOW
offline first, in the existing post-processor; needs lower exposure — the white body currently saturates the sensor
Feed simulation models
measured strain fields calibrating FE material cards
HOW
export the field per frame; compare predicted vs measured maps on the same geometry
Studying fracture behaviour
full-field strain maps around the crack tip; crack propagation speed and direction
HOW
film the necking zone at higher magnification and run the 2D-DIC field around the growing crack — builds directly on the strain-field item above
Correlation, not binary
sub-pixel tracking that keeps the grey levels
HOW
measured gain: 1.1–2.1× less jitter; add as a post-processing option before touching the live loop
Design better grips
printed jaws that seat deterministically — no low-strain slack, no 21 % / 33 % lottery
HOW
size the slack (re-mount one specimen, elastic pulls only), redesign, reprint, re-run — success: one cluster, shorter toe
The upcoming stages are more than features — the 2D-DIC field, fracture behaviour and the simulation coupling are WORTH A THESIS

## Slide 18
06 · OUTLOOK
Interesting numbers — moisture effects & emissions
OUTLOOK
18
MOISTURE → UTS & FAILURE STRAIN
our rig: 4.2 % drop in UTS, Fracture strain UP
the batch whose filament sat unconditioned 8 days longer prints weaker but more ductile — the plasticizing signature:
S7 · (8 days earlier)
batch b–e
offset
UTS
47.82 MPa
45.81 MPa
−4.2 %
εf
0.030
0.05–0.075
↑1.7–2.5×
literature: 8.5 % drop in 24 h
humid filament storage costs PLA strength (150 h: −11 %); ≤65 % RH keeps UTS within 3.7 %.
Hamid et al. 2022 (doi)
Quader et al. 2024 (doi)
CARBON FOOTPRINT – Rough estimate
≈3.2 kg CO₂e
For the WHOLE 6-month project — counted:
▪ the PC work (955 h at 60 W → 57.3 kWh)
▪ the AI compute (195 h of sessions, 0.24 Wh per prompt)
▪ the Mk4S printer (60 specimen prints → 5.0 kWh)
▪ the rig + camera (45 W, all 39 fractures → 0.3 kWh)
— all on Sweden's 0.040 kg CO₂e/kWh grid.
≈239× — thinking over pulling
the AI + the work PC out-emit the test rig itself by that factor — the rig's whole share, all 39 fractures, is ≈13 g. Honest band for the total: 2.6–3.8 kg.

## Slide 19
where it began
where it went
Thank you.
A big thank you to everyone at PPD — for an amazing time here.
Questions — or welcome to break more specimens with me.
Adithya Sivakumar
adithya.sivakumar@ju.se
adithyasvk@outlook.com
19

## Slide 20
APPENDIX
For the curious
APPENDIX
20
A1
Otsu thresholding
how the camera decides what is a marker
A2
PLA vs PETG vs TPU
our curves against what literature expects
A3
Noise, rig against rig
how quiet each instrument is, measured at rest
A4
Camera properties
resolution is a window, exposure sets the frame rate — what was set, and based on what

## Slide 21
APPENDIX · A1
Otsu — how the camera picks its threshold
APPENDIX
21
▪ WHY it exists
▪ every frame must be split into “marker” and “not marker” — that takes one brightness number
▪ WHAT it does
▪ tries all 254 cuts, keeps the one giving the two tidiest groups — lands in the histogram's empty valley
▪ THE CATCH
▪ it follows the SCENE, not the markers — so the rig scores it against fixed values (auto-calibrate) instead of trusting it blindly
Otsu (1979), IEEE Trans. SMC 9(1):62–66 · measured here on 120 frames of specimen S13.

## Slide 22
APPENDIX · A2
Three materials — measured, against expectations
APPENDIX
22
▪ ordering exactly as published
▪ TPU never fractured — it left the frame at 19 % strain
published ranges — PLA: Chacón et al. 2017 · Tymrak et al. 2014 · Prusament PLA TDS
PETG: Durgashyam et al. 2019 · Hsueh et al. 2021 · Prusament PETG TDS — TPU: filament TDS + Hohimer et al. 2020, per the material notes.

## Slide 23
APPENDIX · A3
How quiet is each instrument?
APPENDIX
23
▪ both rigs sit in the tens of microstrain
▪ our floor is set by pixel quantisation: noise ∝ 1/Px₀
▪ predicted 1.78× between our two gauge lengths — measured 1.78×
▪ so the noise is understood, not just small

## Slide 24
APPENDIX · A4
Camera properties — resolution & exposure
APPENDIX
24
RESOLUTION — chosen as a window
▪ the Basler sensor always runs full resolution, 2448 × 2048 px — no binning, no scaling anywhere
▪ what we SET is which pixels to read: ROI [offset 0, 988 · size 2348 × 419], framed with roi_tool.py so both markers sit inside with room to move; rotated 90° → the 420 × 2348 recordings
▪ the ROI is ≈20 % of the sensor (0.98 of 5.0 Mpx) — every frame grabbed, tracked and written is ≈5× lighter, which is what keeps live DIC plus video recording comfortable in real time
▪ the value is a TRAVEL budget: at 2348 px a marker reaches the frame edge at ≈33 % strain — sharpness, scale and noise belong to the optics and exposure, not the ROI
EXPOSURE — sets the frame rate
▪ each frame must be exposed before it can be delivered: at 50 ms per frame the ceiling is 1/0.050 s = 20 fps — the app may ask for 35, the camera delivers 19.9
▪ WHY: the requested rate is a ceiling, not a promise — the sensor cannot start sending a frame until it has finished gathering light for it, so the true rate is min(requested, 1/exposure); readout overhead shaves 20 to 19.9, and below 28.6 ms of exposure the full 35 would return
▪ the 19.9 is MEASURED from frame timestamps on every capture; the video container's declared 35 fps is never believed
▪ the value is chosen by Auto-calibrate: it sweeps ×0.3–×2.8 of the current exposure (15–140 ms), scores each setting on measured marker trackability, and proposes the winner — more light means stronger contrast but fewer frames per second
In short: resolution decides WHERE we look, exposure decides HOW OFTEN — both set once per rig geometry, by tools, not by hand. Full story: the Weekly deck's camera block.