## Slide 1
SMART UTM — a tensile rig that runs, watches and stops its own tests
Jönköping University · School of Engineering — Materials & Manufacturing
The camera measures what a certified extensometer measures.
20 smart features and 6 closed-loop protocols on a desktop rig. On the same frames, the rig's DIC agrees with a Zwick/Roell XT-205 video extensometer to 0.07 % in scale and 0.19 % in fracture strain — and PLA, PETG and TPU land inside their published bands.
18 / 20
smart features
rig-validated (+2 offline)
0.99928
scale vs XT-205 on the
same frames, R² 0.9999964
±1 %
yield and fracture strain
vs the laboratory
12 µε
DIC noise floor, 10 s —
reference instrument 45 µε
1.78× / 1.78×
noise vs pixel span
predicted / measured
3
polymers inside their
published bands
Code, raw CSVs, the 393-slide progress diary and the project report: Software/UTM_PyQt6/ and documentation/ · Every figure and number here is regenerated from the rig CSVs and the laboratory records by the documentation/scripts analysis suite — none of it is transcribed by hand. · Layout follows the #evenbetterposter / Generation-2 billboard format: a plain-language headline, with methods and limitations kept as full panels rather than sidebars.
THE MACHINE
The rig — built in-house, every control loop closed in software
Two steppers (1.85 Nm, 20:1) drive 5 mm lead screws against a fixed base. An ANYLOAD 3 t cell measures force; a Basler camera measures strain optically at 20.8 px/mm from two markers on the 80 mm gauge. The ESP32 runs STEP/DIR/ENABLE only and knows nothing of test protocols — every protocol, guard and measurement here lives in the PyQt6 app, closing the loop at ~11 Hz.
HOW IT WORKS
One control engine
All six protocols are one function called on every load sample (~11 Hz) with load, position and time-matched DIC strain in hand. A protocol only returns a speed — it cannot bypass the shared safety net.
Strain from the specimen
A camera tracks two markers; engineering strain is (L−L₀)/L₀ at 20.8 px/mm. The centroid resolves 0.1 px steps at ±0.02 px.
WHAT THE TESTS SHOWED
Validated against a certified video extensometer — on its own frames
The laboratory's XT-205 video was re-measured by the rig's post-processor: 3153 frames, scale 0.99928, R² 0.9999964, fracture strain 5.21 vs 5.22 %, and a constant offset of -55 µε — 0.069 px on a 1255 px reference — that vanishes from every slope. On separate specimens the two machines agree within 1 % on yield and fracture strain; the -7.7 % UTS gap is pull rate (2.21× the laboratory's), grip seating and batch, not the strain channel.
WHAT IT MEANS
Same frames, same answer → the software is a measurement
Given the XT-205's own frames, the rig's algorithm reproduces the certified record to 0.07 % in scale. Whatever still differs between the two machines is the machine or the specimen, not the strain channel.
The noise is centroid noise → design with pixels
Floor 11.9 µε over 10 s, 26 µε over 900 s, below the reference instrument's 45 µε. It scales as 1/Px₀: 1675 vs 939 px across the gauge predicted 1.78× and measured 1.78×. More pixels across the gauge is the only lever, and it needs no code.
The crosshead is not the specimen → an extensometer is not optional
Only 19–29 % of early crosshead motion reaches the gauge; a crosshead-based curve overstates strain, and understates modulus, by three to five times.
THE SMART FEATURES, SF1 TO SF20
built & rig-validated (18)
built, verified offline (2)
SF1
Live DIC health HUD — 2 / 2
A green/amber/red badge shows marker lock, tracking % and jitter live — so a dropout is seen during the test, not found later in the CSV
SF2
Prepare specimen — 4 → 1
One button tares position, force and DIC and clears the plots
SF3
Recipes / settings — 1 click
Saves and restores an entire setup, including all six protocols' parameters — so repeats are identical by construction, not by memory
SF4
One-click report — 1 click
A one-page PDF and every plot, recomputed from the CSV by the same library the app runs live
SF5
Auto-stop at fracture — 1.2 s
A load-collapse detector halts the motor the moment the specimen breaks
SF6
Constant-strain-rate testing — 0.05 %/s
Closed loop on the specimen's own dε/dt: the crosshead adapted itself 0.10 → 0.05 mm/s, fast in the elastic region, slow through necking
SF7
Stall guard — 0.05 mm / 6 s
Halts a motor that has stopped moving under load
SF8
Release load — ≤ 5 N
Taring at preload makes the readout say 0 N while 300 N is still on the specimen
SF9
Six closed-loop protocols — 6 / 6
cyclic · staircase · relaxation · creep · staircase→FRACTURE · progressive-cyclic→FRACTURE
SF10
Auto-preload — 1.03 ×
Approaches on a 0.2 → 0.1 → 0.02 mm/s schedule and stops 3 % high, because PLA relaxes about 2 % while held
SF11
Auto-metadata link — 3 links
Every save links the CSV, the captured video and the registry row by name and proposes the per-specimen folder — nothing a run produces can be orphaned
SF12
DIC auto-calibrate — 1 click
Exposure and threshold are swept on the live camera and scored on marker count, contrast and centroid jitter; the winner is applied with the score shown
SF13
Guided wizard — 9 / 9
Connect → camera → data → mode → auto-calibrate → preload → Px₀ → prepare → run → save
SF14
Test registry — 36 runs
Each run indexed with whatever its loading path supports — UTS, ε_f, E, σ_y, toughness — plus the force anchor that turns a tared reading into a true load
SF15
Dead-DIC guard + hard backstops — 3.0×
A controller ramping blind on a stale camera reading spikes the force
SF16
Live Px₀ overlay — Δ = ε
The frozen reference pair and the live pair are drawn on the feed with a caption whose Δ agrees with the strain readout — the operator sees the measurement, not only a number
SF17
Video + image capture — ≈47 µs
PNG stills plus three AVI views (raw, contrast-boosted, adaptive speckle) recorded with the run, at a measured cost that never touches the detection budget
SF18
DIC post-processing — 0.99928
Measures strain from ANY recorded video with the rig's own pixel-to-strain rule
SF19
Noise capture — 26 µε
Records the instrument at rest and separates offset and drift (correctable) from the residual sd — the uncertainty, which does not subtract
SF20
Manual blob selection — 9 / 9
The operator clicks the two markers on a frozen frame and tracking follows the picks with the shape gates off — keeps measuring where automatic detection is deliberately killed