DIC STUDY GUIDE

Digital Image Correlation

From Beginner to Practitioner

Based on: iDICs Good Practices Guide for DIC, Edition 2 (October 2025)

International Digital Image Correlation Society (iDICs)

# 1. Document Summary

This study guide is based on the iDICs Good Practices Guide for Digital Image Correlation, Edition 2 (October 2025). The guide is a 166-page comprehensive reference produced by the International Digital Image Correlation Society (iDICs), involving contributions from academia, government, and industry experts including representatives from all major commercial DIC software packages.

The guide covers everything from fundamental DIC principles to reporting requirements, and is structured as follows:

| Chapter 1 | Introduction — aims, basic principles, and scope of DIC |
| --- | --- |
| Chapter 2 | Design of DIC Measurements — hardware, optics, patterns, environmental factors |
| Chapter 3 | Preparation for Measurements — pre-calibration routine, calibration procedure |
| Chapter 4 | Execution of the Test — image acquisition workflow |
| Chapter 5 | DIC Analysis — image processing, subset parameters, uncertainty, VSG study |
| Chapter 6 | Reporting Requirements — what must be documented in every DIC test |
| Chapter 7 | Glossary and Acronyms — definitions of all terms |
| Appendices A–D | FOV/lens calculations, depth-of-field, strain calculation methods, global DIC |

| NOTE | The guide is not a rigid standard but a good practices reference. However, some guidelines are described as crucial while others are recommended — both are important for trustworthy results. |
| --- | --- |

# 2. What is DIC? — Core Principles

## 2.1 Basic Concept

Digital Image Correlation (DIC) is an optical, non-contact, full-field measurement technique. It works by tracking the movement of a random pattern (typically a speckle pattern sprayed on the specimen surface) between images taken before and after deformation. By comparing corresponding regions (subsets) in the reference and deformed images, the software computes displacement and strain fields across the entire surface.

| TIP | DIC gives you full-field displacement and strain maps — unlike a strain gauge which only measures at one point. This is its key advantage. |
| --- | --- |

## 2.2 The Two Types of DIC

| 2D-DIC (Start Here) | Stereo-DIC (2D-DIC = next step) |
| --- | --- |
| One camera Measures in-plane (X, Y) displacements only Assumes specimen stays flat (planar) throughout the test Any out-of-plane motion causes false strain Simpler setup, good starting point | Two cameras at a stereo-angle Measures 3D surface coordinates (X, Y, Z) and full 3D displacements Handles out-of-plane motion correctly Requires stereo calibration More complex but more robust |

| TIP | For your first DIC test, use 2D-DIC. The planar assumption is acceptable for flat tensile specimens when the camera is perpendicular to the specimen surface. |
| --- | --- |

## 2.3 Key Terminology You Must Know

| QOI | Quantity-of-Interest — what you want to measure (displacement, strain, etc.) |
| --- | --- |
| ROI | Region-of-Interest — the area on the specimen you want to analyze |
| FOV | Field-of-View — the physical area captured in the camera image |
| SOD | Stand-Off Distance — distance from camera to specimen |
| DOF | Depth-of-Field — the range of distances along the optical axis that appear in focus |
| Subset | A small rectangular region of the image (e.g. 21x21 pixels) used for correlation |
| Step Size | The spacing between adjacent subsets/data points in the image |
| VSG | Virtual Strain Gauge — the physical size over which strain is averaged in DIC |
| Noise-floor | The minimum detectable signal level; anything below this is noise |
| Image Scale | Conversion factor from pixels to physical units (e.g. px/mm) |
| Calibration | Process of determining camera intrinsic/extrinsic parameters using known targets |

## 2.4 How DIC Works — Step by Step

Understanding the DIC workflow is essential before designing your test:

- Apply a random speckle pattern to the specimen surface

- Set up and calibrate your camera/lens system

- Acquire a reference image before loading/deformation

- Load the specimen and acquire images during/after deformation

- DIC software divides the ROI into subsets and correlates them between reference and deformed images

- Displacement fields (U, V) are computed at each data point

- Strain fields (exx, eyy, exy) are calculated from spatial derivatives of displacement

# 3. Setting Up Your DIC Test Rig

## 3.1 Hardware Components Required

For a basic 1D/2D-DIC setup on a mechanical testing machine, you need:

- Camera: Machine-vision monochromatic camera (global shutter recommended)

- Lens: Fixed focal-length lens with lockable focus, aperture, and zoom rings

- Mounting: Rigid camera mounting structure (tripod, optical rail, or custom rig)

- Lighting: External lighting (LED panels preferred)

- Computer: Computer with DIC software and sufficient storage

- Calibration target: Calibration target (dot grid or checkerboard, approximately FOV size)

## 3.2 Camera Selection — Key Factors

| Sensor type | Monochromatic (grayscale) — color cameras can be used but require additional processing |
| --- | --- |
| Shutter type | Global shutter preferred (reads all pixels simultaneously). Avoid rolling shutters for DIC |
| Auto functions | DISABLE all auto-focus, auto-aperture, auto-exposure — these will ruin DIC measurements |
| Bit depth | 12-bit or higher gives better dynamic range than 8-bit |
| Resolution | Higher Mpx = better spatial resolution; common choices: 1.3, 5, 12, 20 Mpx |
| Frame rate | Must be high enough to capture your fastest event; governed by test speed |

| CAUTION | The most important thing for camera selection: disable ALL automatic adjustments. Once set, nothing should change between your reference image and deformed images. |
| --- | --- |

## 3.3 Lens Selection — Key Factors

| Focal length | Determines FOV for a given SOD. Longer focal length = smaller FOV but longer SOD |
| --- | --- |
| Fixed vs zoom | Fixed focal length preferred — zoom lenses have higher distortions and may shift accidentally |
| Lockable rings | Choose lenses where focus ring, aperture ring, and zoom ring can be locked |
| Telecentric lens | Recommended for 2D-DIC to minimize errors from small out-of-plane motions |
| Distortions | Low distortion lenses preferred — calibration can correct some distortions but not all |

| TIP | For 2D-DIC: use as long a focal length as possible and maximize the Stand-Off Distance (SOD). This minimizes the false strain caused by any out-of-plane motion. |
| --- | --- |

## 3.4 Camera Mounting — Critical Requirements

The mounting system is one of the most overlooked aspects of DIC. Poor mounting introduces vibrations and camera drift that appear as false strain. Requirements:

- Rigid and stable — no vibrations, no thermal drift, no creep

- Sufficient degrees of freedom to precisely position and align the camera(s)

- Lockable adjustments — once set, nothing should move during the test

- Decoupled from the test machine — camera should not feel vibrations from the load frame

- Thermally stable — consider heat expansion of metals in long tests

| TIP | For a simple test: a sturdy tripod with a ball head is acceptable, but a rigid optical rail or aluminum extrusion rig clamped to a vibration-isolated table is much better. |
| --- | --- |

## 3.5 Lighting — Setup and Configuration

Lighting quality directly impacts DIC accuracy. Poor lighting is a common cause of DIC measurement errors.

### 3.5.1 Lighting Type

- Best choice: LED panels: recommended — stable output, low heat, no flicker

- Acceptable: Fluorescent: some flicker at 50/60 Hz can cause intensity changes between frames — use with caution

- Avoid: Halogen/incandescent: high heat, thermal drift — avoid if possible

- Special use: Strobe/flash: can freeze motion for high-speed tests, requires synchronization

### 3.5.2 Lighting Geometry

- Use diffuse lighting (large soft-boxes or ring lights) to avoid specular reflections (glare)

- Position lights at an angle to the specimen surface — avoid lighting directly along the camera optical axis

- For 2D-DIC: symmetric lighting from both sides of the specimen is ideal

- Consider cross-polarized lighting to eliminate glare on shiny specimens

- Ensure light coverage is uniform across the entire FOV

### 3.5.3 Critical Lighting Rules

| CAUTION | Once you set lighting intensity and position, DO NOT change it during the test. Changing lighting changes image brightness, which causes false displacements in DIC. Lock everything down. |
| --- | --- |

- Let LED lights warm up (5–30 min) before testing to reach stable output

- Avoid light sources that heat the specimen — heat waves cause optical distortions in DIC results

- Use fans to cool hot lights, but mount fans so vibrations don't reach the camera

## 3.6 2D-DIC vs Stereo-DIC Setup Checklist

| Requirement | 2D-DIC | Stereo-DIC |
| --- | --- | --- |
| Number of cameras | 1 | 2 (stereo pair) |
| Camera alignment | Perpendicular to specimen, optical axis normal to specimen surface | Angled at stereo-angle to each other |
| Stereo-angle | N/A | 15–35 degrees typical |
| Calibration type | Scale calibration (and lens distortion correction if needed) | Full 3D stereo calibration required |
| Camera synchronization | Not required | CRITICAL — cameras must fire simultaneously |
| Out-of-plane sensitivity | NOT measured — errors if specimen moves out-of-plane | Fully measured in all 3 axes |
| Specimen flatness required | YES — must remain flat and perpendicular to camera | Not required |
| Setup complexity | Simple | More complex |

# 4. Key Factors and How Each Affects Results

## 4.1 Speckle Pattern — The Foundation of DIC

The speckle pattern is the single most important element of a DIC test. A poor pattern will cause poor or failed correlation regardless of how good your camera and software are.

### 4.1.1 Pattern Feature Size

| Recommended size | 3–5 pixels per speckle feature in the image |
| --- | --- |
| Too small (<3 px) | Aliasing errors — features cannot be resolved correctly by the sensor |
| Too large (>5 px) | Requires larger subsets — degrades spatial resolution of displacement/strain fields |
| Physical size | Depends on image scale (px/mm). Calculate: L_physical = L_image / S (where S = px/mm) |
| Verification | Zoom into your image and count pixels across a feature by eye — aim for 3–5 px |

| TIP | Example: If your image scale is 10 px/mm and you want 4 px features, your physical speckle diameter should be 4/10 = 0.4 mm. |
| --- | --- |

### 4.1.2 Pattern Characteristics — What Makes a Good Pattern

- Random: speckles should be randomly distributed — no regular grid or repeating patterns

- Matte finish: use flat/matte paint — glossy speckles cause specular reflections (glare) that change between images

- Approximately 50% coverage: roughly half black, half white across the pattern area

- Uniform feature size: consistent speckle diameter across the ROI

- High contrast: dark speckles on a light background (or vice versa) — avoid low-contrast patterns

- Good bonding: pattern must not peel, crack, or debond during the test — this is critical

### 4.1.3 Common Patterning Methods

| Spray paint (rattle can) | Most common. White base coat + black overspray. Control nozzle distance and duration. Best for specimens > ~20 mm |
| --- | --- |
| Airbrush | Better control of feature size. Good for small specimens |
| Stencil / stamp | Avoid — can create repeated or regular patterns which confuse the correlation algorithm |
| Toner powder | Fine features possible for microscale DIC |
| Natural texture | If the specimen surface has natural contrast (e.g. woven fabric, concrete), evaluate first — may not need applied pattern |
| Etching / abrasion | Removes material — check that specimen properties are not affected |

| TIP | For your first test: white base coat (spray paint), let dry, then light black overspray at arm's length for ~0.5 seconds. Check the image — aim for 3–5 px features. |
| --- | --- |

### 4.1.4 Pattern Application Rules

- Clean the specimen surface before painting (acetone wipe for metals removes grease)

- Let the pattern set before testing, but test before full cure if ductile paint is needed for large deformation

- Apply fiducial marks (two dots a known distance apart) to verify the image scale

- Do not touch the pattern after application — oil from fingers can cause debonding

| CAUTION | Inspect the pattern after the test! Look for cracking, debonding, or fading. If the pattern failed during the test, your DIC results in that region are invalid. |
| --- | --- |

## 4.2 Field-of-View (FOV) and Image Scale

The FOV defines the physical area captured by your camera. The image scale (px/mm) is the conversion factor between pixels and real-world distances. These affect spatial resolution, pattern feature size, and noise-floor.

| Rule | ROI should nearly fill the FOV — maximizes spatial resolution while keeping specimen in view |
| --- | --- |
| Image scale formula | S (px/mm) = image width (px) / FOV width (mm) |
| Effect on pattern | Smaller FOV = smaller physical speckle size needed to stay at 3–5 px |
| Effect on strain noise | Smaller FOV = smaller VSG in mm = potentially higher strain noise |
| Effect on gradients | Smaller FOV = can resolve finer spatial strain gradients |

| NOTE | FOV trade-off: A smaller FOV gives better spatial resolution but may increase strain noise. A larger FOV covers more area but may miss fine strain gradients. |
| --- | --- |

## 4.3 Aperture (f-number)

The aperture controls how much light enters the lens and directly affects Depth-of-Field (DOF) and image brightness.

| Small f-number (wide open, e.g. f/2.8) | More light, shallow DOF — easier to focus precisely, less tolerant of specimen misalignment |
| --- | --- |
| Large f-number (stopped down, e.g. f/16) | Less light, deep DOF — keeps more in focus, requires brighter lighting |
| For 2D-DIC | DOF is not critical (flat specimen), but sufficient DOF helps tolerate minor misalignment |
| For stereo-DIC | DOF must cover the full specimen depth as seen from the stereo-angle |
| Recommended approach | Set DOF first (based on test needs), then adjust lighting to compensate for exposure |

| CAUTION | Once aperture is set and calibration is done — DO NOT change the aperture. Changing it changes the focus and DOF, invalidating your calibration. |
| --- | --- |

## 4.4 Exposure Time

Exposure time controls image brightness and motion blur. There is a fundamental trade-off.

| Motion blur rule | Motion blur < 0.5 pixels during exposure. Estimate: speed (px/s) x exposure time (s) < 0.5 px |
| --- | --- |
| Too long | Motion blur smears speckles, degrades correlation, produces wrong displacements |
| Too short | Image too dark — reduces image quality and increases noise in DIC results |
| Exposure limit | Exposure time < 1/frame rate (cannot be longer than the inter-frame period) |
| Typical range | For quasi-static tests: 1–50 ms. For high-speed: microseconds. |
| Fix for too-short exposure | Add more light (brighter LEDs, more sources) then decrease exposure time |

| TIP | Gray level target: aim for 50–80% of dynamic range (not saturated, not too dark). For 8-bit camera: 128–200 gray levels. For 12-bit: 2000–3200 counts. |
| --- | --- |

## 4.5 Frame Rate and Image Acquisition Rate

These control the temporal resolution of your DIC measurements.

| Key question | How much does the specimen move between images? Displacement should be < ~half the subset size |
| --- | --- |
| Quasi-static tests | 1–10 Hz typically sufficient; capture enough images to characterize the full loading curve |
| Dynamic/impact tests | Must capture all relevant deformation events — consult Nyquist theorem for cyclic loading |
| Cyclic loading rule | Image acquisition rate must be at least 2x the loading frequency (Nyquist), preferably much higher |
| Maximum displacement per frame | Rule of thumb: < half the subset size in pixels (e.g. < 12.5 px for a 25 px subset) |
| Data management | DIC images generate large files — plan storage capacity before testing |

## 4.6 Subset Size and Step Size

These are the primary DIC software parameters. They directly control spatial resolution, noise-floor, and the size of the Virtual Strain Gauge.

### 4.6.1 Subset Size

| Definition | Side length of the square image region used for correlation (e.g. 21x21 pixels = 21 px subset) |
| --- | --- |
| Typical range | 15–71 pixels; starting point: 31–41 px for a well-made pattern |
| Minimum rule | Subset must contain at least 3–5 unique speckle features |
| Too small | Not enough pattern information — correlation fails or is unreliable (high noise) |
| Too large | Smooths over local deformation gradients — misses strain concentrations, bias error increases |
| Effect on VSG | Larger subset = larger VSG = smoother strain field but lower spatial resolution |
| Pattern dependency | Coarser pattern requires larger subset; finer pattern allows smaller subset |

### 4.6.2 Step Size

| Definition | Spacing between adjacent data points (subset centers) in the image |
| --- | --- |
| Typical range | 1/3 to 1/5 of the subset size is a common starting point (e.g. 7–11 px for a 31 px subset) |
| Smaller step size | More data points, better ability to detect peak strains, more computation time |
| Larger step size | Fewer data points, faster processing, may miss strain peaks |
| Effect on VSG | Step size affects VSG size — see VSG formula below |
| For strain peaks | Use a small step size (2–5 px) to avoid missing the peak due to undersampling |

### 4.6.3 Virtual Strain Gauge (VSG) Size — The Key Metric

The VSG is the physical gauge length over which strain is effectively averaged in DIC. It is the DIC equivalent of a physical strain gauge.

VSG (px) approximately = (subset size) + (strain window - 1) x (step size)

| Smaller VSG | Higher spatial resolution, captures strain peaks better, but higher noise-floor |
| --- | --- |
| Larger VSG | Smoother strain field, lower noise, but averages over more area (misses peaks) |
| VSG study | Always perform a VSG study: vary the VSG size and plot peak strain vs. noise-floor to find the optimal balance |
| Converged result | When peak strain stops increasing as VSG decreases — you have resolved the peak |

## 4.7 Depth-of-Field (DOF)

DOF is the range along the optical axis where objects appear acceptably sharp.

| Controls | Primarily set by the lens aperture (f-number). Smaller aperture (larger f-number) = deeper DOF |
| --- | --- |
| For 2D-DIC | DOF not critical for a flat specimen, but having enough DOF reduces sensitivity to minor misalignment |
| For stereo-DIC | DOF must be sufficient to keep the specimen in focus across the entire stereo viewing volume |
| Focus setting | Open aperture wide to limit DOF, find the sharpest focus, then stop down to desired aperture |
| DOF calculation | Appendix B of the guide gives formulas. Can also be measured with a calibrated target at 45 degrees |

## 4.8 Noise-Floor

The noise-floor defines the minimum measurable signal. Anything smaller than the noise-floor is indistinguishable from random measurement error.

| Measurement | Acquire several images of the static, unloaded specimen and compute the standard deviation of displacement/strain results |
| --- | --- |
| Displacement noise | Typically 0.01–0.05 px for good DIC setups |
| Strain noise | Typically 50–500 microstrain depending on VSG size, pattern quality, lighting |
| Trade-off | Smaller VSG = higher strain noise but better spatial resolution. Larger VSG = lower noise but smoothed strains |
| Evaluation timing | Evaluate noise-floor before the test during pre-calibration review, and report it with results |

| NOTE | Always evaluate and report your noise-floor. This tells readers (and yourself) what the minimum meaningful strain measurement from your system is. |
| --- | --- |

# 5. Calibration

## 5.1 Purpose of Calibration

Calibration determines the intrinsic parameters of each camera (image scale, focal length, lens distortions) and, for stereo-DIC, the extrinsic parameters (position and orientation of cameras relative to each other and to the specimen).

- 2D-DIC calibration: primarily to determine image scale and correct lens distortions

- Stereo-DIC calibration: full 3D calibration using a calibration target in multiple orientations

| NOTE | For 2D-DIC: if lens distortions are negligible compared to your noise-floor, you may skip calibration and just use the image scale from fiducial marks. Verify this by translating your specimen in-plane and checking for false strains. |
| --- | --- |

## 5.2 Calibration Target

| Target type | Dot grid (circles) or checkerboard — must match your DIC software |
| --- | --- |
| Target size | Should be approximately the same size as your FOV (at minimum, half the FOV) |
| Feature spacing accuracy | Should be accurate to within 0.1 pixels — use metrologically traceable targets |
| Material | Rigid, flat plate — do not use paper targets without a flat backing |
| Printed paper targets | Acceptable if glued to a rigid plate and feature spacing is verified |

## 5.3 Pre-Calibration Checklist

Before calibrating, complete these steps in order:

- Warm up cameras: run camera at test frame rate for 30–60 minutes until thermal equilibrium. Camera heat causes thermal expansion and drift in DIC results if not stabilized

- Clean optics: check lens and camera sensor for dust/dirt. Image a white background and look for blurred spots

- Mount specimen in the test machine and position camera to achieve the desired FOV

- Set focus: open aperture wide, zoom in on live image, adjust focus ring until speckles are sharp. Then stop down to test aperture

- Set lighting and exposure: aim for 50–80% of dynamic range in gray level. Check for uniform illumination

- Lock all adjustable components: focus ring, aperture, zoom, camera position — nothing should move during the test

- Review images: check for glare, out-of-focus areas, non-uniform lighting, dirt on lens, pattern issues

- Correlate static images: run DIC on sequential static images. Any non-zero displacement indicates a problem

## 5.4 Calibration Procedure for 2D-DIC

- Place calibration target where the specimen will be during testing

- Adjust lighting to be appropriate for the calibration target (may differ from specimen lighting)

- Acquire calibration images (follow software-specific requirements — typically 10–50 images)

- Process calibration in DIC software — review calibration residuals/score

- Note: aperture and focus must NOT change between calibration and the test

## 5.5 Calibration Procedure for Stereo-DIC

For stereo-DIC, the calibration target must be placed in the calibration volume (the space where the specimen will be) and imaged in many orientations:

- Rotate target about X axis (tilt forward/back)

- Rotate target about Y axis (tilt left/right)

- Rotate target about Z axis (in-plane rotation)

- Plunge target towards and away from cameras along Z axis

- Translate target horizontally and vertically if target is smaller than FOV

- Rotate target 90 degrees about Z and repeat all above

- Perform combinations of the above movements

| Number of images | 8 up to 50–100 depending on software; check your software manual |
| --- | --- |
| 3D calibration targets | May require fewer images than flat 2D targets |
| Calibration score | Low residual = better calibration. Consult your software for acceptable thresholds |
| After calibration | Do NOT move the cameras relative to each other. They must remain rigidly fixed |

| CAUTION | For stereo-DIC: if you bump or move the cameras relative to each other after calibration, you MUST re-calibrate before the test is valid. |
| --- | --- |

# 6. Performing the DIC Test

## 6.1 Pre-Test Checklist

- Camera warm-up complete (30–60+ min)

- Pattern applied and inspected — features are 3–5 px in image

- Calibration complete and verified

- Lighting stable and locked

- Aperture, focus, zoom locked

- Reference image strategy defined (typically: image at zero load, before test starts)

- Synchronization verified (for stereo-DIC or multi-instrument tests)

- Storage space confirmed — estimate image count x image file size

- Test procedure reviewed

## 6.2 Acquiring the Reference Image

The reference image is the baseline against which all deformed images are compared. Getting this right is critical.

| Standard approach | Acquire reference image before any loading, at zero force |
| --- | --- |
| Timing | Take multiple images at rest — use the one with the best quality |
| No motion allowed | Any motion between the reference image and subsequent images introduces errors |
| Incremental approach | Some advanced tests use multiple reference images updated during the test — but start with a single reference |

## 6.3 During the Test

- Acquire images at the planned image acquisition rate

- Monitor image quality in real-time if possible — watch for overexposure, blur, or correlation failure indicators

- Do not touch, bump, or adjust the camera, lens, or mounting at any point

- Log any anomalies (lighting changes, vibrations, etc.) with timestamps for post-test review

- For synchronized tests: verify external signals (force, extensometer) are being recorded simultaneously

## 6.4 After the Test

- Inspect the specimen and pattern — document any cracking, debonding, or pattern failure

- Acquire post-test static images — useful for comparison and noise evaluation

- Save all raw images in a lossless format (TIFF preferred, avoid JPEG for DIC)

- Back up raw images before processing

# 7. DIC Analysis — Key Software Parameters

## 7.1 Analysis Workflow

- Load reference and deformed images into DIC software

- Define the Region-of-Interest (ROI) in the image

- Set subset size and step size

- Choose correlation criterion (ZNSSD recommended — robust to lighting changes)

- Run correlation to compute displacement fields

- Set strain window and strain tensor type

- Compute strain fields

- Evaluate noise-floor and perform VSG study to optimize parameters

## 7.2 Important Analysis Parameters

| Matching criterion | ZNSSD (Zero-mean Normalized Sum of Squared Differences) — most robust to lighting changes. CC and SSD also used |
| --- | --- |
| Subset shape function | Affine (first-order) sufficient for most tests. Quadratic for high-curvature deformations |
| Interpolant | Bi-cubic spline is typical — needed for sub-pixel accuracy |
| Strain tensor | Engineering strain for small deformation; Green-Lagrange or Hencky (true) strain for large deformation |
| Image filtering | Optional Gaussian pre-filter can reduce camera noise, but may slightly smear the pattern |
| Initial guess | Required if displacement between frames is large (>subset size/2); usually automatic in software |

## 7.3 VSG Study — Finding Your Optimal Parameters

The VSG study is how you determine the best trade-off between noise and spatial resolution for your specific test. Always perform this.

- Select 10 static images (no loading) and analyze with various VSG sizes

- For each VSG size, compute the noise-floor (standard deviation of strain in static images)

- Select the image with the highest strain gradient from your loaded images

- Extract a line cut through the highest strain gradient region for each VSG size

- Plot: peak strain vs VSG size and noise-floor vs VSG size

- Choose the smallest VSG where peak strain has converged (stopped changing) and noise is acceptable

| CAUTION | If peak strain keeps increasing as VSG decreases and never converges: the true peak is not resolved. You need a smaller FOV or higher resolution camera. You can only report a lower bound on the true peak strain. |
| --- | --- |

# 8. Starting with 1D DIC — Your First Steps

## 8.1 What is 1D DIC?

Strictly speaking, DIC is always a 2D or 3D analysis (it tracks a 2D image). What people often call '1D DIC' is extracting a 1D measurement from 2D DIC — for example, extracting a single axial strain value (like a virtual extensometer) from the full-field DIC result. This is the simplest use of DIC and an excellent starting point.

## 8.2 Simple 1D Use Cases

- Virtual extensometer: measure average displacement between two points on the specimen (like an extensometer)

- Average axial strain in a gauge section: extract the mean strain over the gauge length from DIC

- Strain along a line: plot strain as a function of position along the specimen axis

## 8.3 Minimal Setup for Your First 1D DIC Test

| Specimen | Flat tensile coupon — the simplest case for 2D-DIC (planar assumption is easily met) |
| --- | --- |
| Camera | Single camera, mounted perpendicular to specimen face |
| Lens | Fixed focal length, manual focus, lockable aperture |
| FOV | Capture the full gauge length with some margin |
| Pattern | White base coat + black overspray. Aim for 3–5 px speckle size |
| Lighting | Two LED panels at ~45 degrees to the specimen, symmetric |
| Calibration | For a simple first test: use fiducial marks a known distance apart to determine image scale. Verify lens distortions are negligible |
| Subset size | Start with 31–41 px and adjust based on VSG study |
| Step size | Start with 1/3 to 1/5 of subset size |
| Reference image | One image at zero load before the test begins |

## 8.4 Key Things to Verify Before Your First Test

- Correlate two sequential static images — displacement should be < 0.05 px (noise-floor)

- Check that image scale is correct by measuring known fiducial marks

- Verify that 50–80% gray level range is used in the image (not saturated, not too dark)

- Confirm all adjustable components are locked

- Confirm pattern features are 3–5 px in the image

## 8.5 Progressing to Full 2D-DIC

Once you are comfortable with 1D extraction from DIC, the step to using the full 2D displacement and strain fields is small — you are already running the same analysis, just using more of the output data. The key additions are:

- Perform a proper VSG study to optimize your parameters

- Evaluate and report the noise-floor for all three strain components (exx, eyy, exy)

- Consider whether lens distortion correction is needed for your application

- Use full-field strain contours instead of line-extracted values

# 9. What You Must Report (Reporting Requirements)

## 9.1 Required Hardware Parameters

- Camera manufacturer and model, and image size (px x px)

- Lens manufacturer, model, focal length, and aperture used

- Field-of-View (FOV) in physical units

- Image scale (px/mm or mm/px) and how it was determined

- Lighting type and configuration

- Patterning method used

## 9.2 Required Analysis Parameters

- DIC software: manufacturer, package name, version number

- Image filtering applied (if any)

- Reference image definition (e.g. single image at zero force)

- Interpolant type (e.g. bi-cubic spline)

- Matching criterion (e.g. ZNSSD)

- Subset size in pixels AND in physical units (mm)

- Step size in pixels AND in physical units (mm)

- Subset shape function (e.g. affine)

- Strain window size (data points)

- Virtual Strain Gauge (VSG) size in pixels AND in physical units

- Strain tensor type (e.g. engineering, Green-Lagrange, Hencky)

- Noise-floor of all QOIs — method used to compute it must be stated

| NOTE | For 2D-DIC: also report any bias errors due to out-of-plane motion, if estimated. |
| --- | --- |

# 10. Quick Reference — Key Numbers to Remember

| Parameter | Target Value | Why It Matters |
| --- | --- | --- |
| Speckle feature size | 3–5 pixels | Below 3 px: aliasing. Above 5 px: large subsets needed |
| Pattern coverage | ~50% dark/light | Good contrast and uniqueness for correlation |
| Image gray level | 50–80% of max | Not saturated, not too dark |
| Displacement noise-floor | < 0.01–0.05 px | Minimum detectable displacement signal |
| Subset size (starting) | 31–41 px | Contains 3–5 speckles; adjust with VSG study |
| Step size (starting) | 1/3 to 1/5 of subset | Balance data density and computation time |
| Motion blur limit | < 0.5 px per exposure | Blurring degrades correlation |
| Stereo-angle | 15–35 degrees | Balance in-plane vs out-of-plane accuracy |
| Calibration target size | ~Same size as FOV (min 1/2 FOV) | Must cover the full calibration volume |
| Camera warm-up | 30–60+ minutes | Prevents thermal drift errors |
| SOD (2D-DIC false strain) | Longer SOD = smaller false strain | At 250 mm SOD: 1 mm OOP motion = 0.4% false strain |
| Calibration images | 8–100 depending on software | More positions/orientations = better calibration |

## 10.1 False Strain from Out-of-Plane Motion in 2D-DIC

This is the most common error in 2D-DIC and is directly related to the Stand-Off Distance (SOD):

| SOD = 250 mm, OOP motion = 1 mm | False strain = 0.4% |
| --- | --- |
| SOD = 500 mm, OOP motion = 1 mm | False strain = 0.2% |
| SOD = 1000 mm, OOP motion = 1 mm | False strain = 0.1% |

| CAUTION | Always estimate the expected out-of-plane motion in your test and calculate the false strain. If it is comparable to or larger than your real strain signal or noise-floor, you should use stereo-DIC instead. |
| --- | --- |

Good luck with your DIC tests!

Reference: iDICs Good Practices Guide for DIC, Ed. 2, Oct. 2025 | www.idics.org